## Slide 1: The LV600 The Spiritual Successor to the Tektronix LV500
Norm Gifford
Daniel Khoury
Project Advisor: Ken Stevens

## Slide 2: Motivation
CS 6712 requires chip testing, but…
  LV500 is unreliable
    Calibration fails
    Several sectors are dead
  Verigy 93000 is not operational
    HDD is dead
    Support expired
Potential legacy
Fun, challenging project
Notes: Norm

## Slide 3: Senior Project (Fall 2014)
Notes: Norm

## Slide 4: Senior Project (Fall 2014)
Notes: Norm
Slow system (huge latency between sending test vector and getting output vector)

## Slide 5: Senior Project (Fall 2014)
What worked well:
  Functional testing!
  Unfixed power and ground locations
  Manual schmooing
Areas of improvement:
  Real-time testing
  Autoschmooing
  Support for multiple templates
  Interface and setup
  Configuration options for test cycles
  Accurate timing information
Notes: Norm

## Slide 6: Master’s Project: Initial Plan
Redo our senior project, but…
  Simplify the user interface and setup
  Replace the FPGA logic with a MOSIS-fabricated chip
  Try to automate schmooing
  Equilibrium tracing (for accurate timing info)
Issues:
  Constrained by MOSIS fabrication deadlines
  Lack of versatility
Notes: Daniel

## Slide 7: Master’s Project: Revised Plan
Complete revamp of our senior project design
Main goals:
  Simpler interface and setup
  Real-time testing
  Timing information (via equilibrium tracing)
  Automated schmooing
  More test cycle configurations
Fall 2015 semester dedicated to planning and prototyping
Spring 2016 semester dedicated to implementation
Notes: Daniel

## Slide 8: Simpler Interface and Setup
BeagleBone Black web server
  Web form to fill out and submit tests
  Commands FPGA via RX/TX lines
Numato Saturn FPGA
  Configuration programs flashes Verilog onto the board
Notes: Norm

## Slide 9: Real-Time Testing
All bit vectors stored in block RAM for fast read access
Notes: Norm

## Slide 10: Automated Schmooing
Notes: Daniel

## Slide 11: Automated Schmooing
Notes: Daniel

## Slide 12: Test Cycle Configurations: Original Idea
Notes: Daniel

## Slide 13: Test Cycle Configurations: Later Idea
Key insight: No need to create the actual test cycle waveforms inside the FPGA
Instead, count clock edges
  User specifies leading edges, trailing edge, cycle length as multiples of 10 ns
Notes: Daniel

## Slide 14: PCBs
Three PCBs:
  Numato shield
  Communication board
  DUT board
Boards designed with versatility in mind
Notes: Norm

## Slide 15: PCBs: Numato Shield
Four-layer board
Manually routed
Equilibrium tracing
  Surprisingly tedious and difficult
  Tolerance: 200 mil
Notes: Norm

## Slide 16: PCBs: Communication Board
Two-layer board
Autorouted
Core functionality:
  Connects BBB to FPGA via RX/TX lines
  Connects BBB to schmooing circuitry
  Connects FPGA to DUT board
  Contains voltage translators (between FPGA and DUT)
Notes: Norm

## Slide 17: PCBs: DUT Board
Four-layer board
Autorouted
  Trickiest part was component placement
  100 minutes to route!
Contains:
  ZIF socket (thank you Erik!)
  Four counters
  Eight SRAM blocks
  16 voltage translators
  16 parallel->serial shift registers
Notes: Norm

## Slide 18: Software Flow
User fills out and submits web form
BeagleBone Black (BBB) transmits all bit vectors and configuration data to FPGA
BBB instructs FPGA to execute tests
FPGA signals BBB when tests are complete
BBB commands FPGA to send back output vectors
BBB displays test results
Notes: Norm

## Slide 19: Software Flow (Continued)
Mostly PHP and HTML
  PHP provides exec function, can call C code that parses web form input
  Output stream (printf statements) passed back to web form
AJAX used to dynamically add/remove vectors
(Online)
C code consists of UART and SPI libraries
  UART for FPGA communication
  SPI to control digital pots in power circuitry
Notes: Norm

## Slide 20: Hardware Flow
Handshaking protocol between BBB and FPGA
Commands BBB can send to FPGA:
  Store input vector
  Store cycle vector
  Store template vector
  Store force-format vector
  Store cycle configuration data
  Execute tests
  Send back output vector
Notes: Daniel

## Slide 21: Hardware Flow
All vectors are 128-bit wide
Up to four templates
Total number of input vectors: 244
  Good enough for most chips
Notes: Daniel

## Slide 22: Hardware Flow (Continued)
FPGA responds to BBB commands as follows
Store input/cycle/template/force-format vector:
  Save vector in block RAM
  If input vector, track how many vectors have been written
Store cycle configuration data:
  Load counters with delays for leading edges, trailing edge, cycle length
Notes: Daniel

## Slide 23: Hardware Flow (Continued)
Execute tests:
  While there are more input vectors to read:
    Beginning of cycle: Fetch next input vector
    Leading edge: Apply inputs, enable SRAM blocks (write-mode)
    Trailing edge: Disable SRAM blocks
    End of cycle: Increment counter
      If the template associated with the new input vector has not changed, load the fetched input vector
      Else pause real-time testing, load the template, cycle, and force-format vectors associated with the new template, then load the fetched input vector and resume real-time testing
  Reset the counter
Notes: Daniel

## Slide 24: Design Challenges: PCBs
Communication board failed twice
  First time: Short between 3V3 and GND (manufacturing error?) 
  Second time: Power and ground reversed on barrel jacks 
  Third time is the charm! 
    Should have ready mid-May.
Autorouting the DUT board proved difficult
  Component placement is key!
Notes: Norm

## Slide 25: Design Challenges: Verilog
Initial Verilog design included R0, R1, DNRZ_L, DNRZ_T force formats (2-bit encoding)
  Design required 108% of available LUTs
Removed DNRZ_T and R1 force formats (1-bit encoding)
  Design required 70% of available LUTs, but 115% of FPGA area
Removed ability to specify two different test cycles, replaced with one test cycle but two leading edges
  Design now utilizes 40% of available LUTs, 70% of available area
Some timing-sensitive FPGA outputs (e.g. counter clock) are glitchy
  Capacitors may smooth out glitches – need to investigate
Notes: Daniel

## Slide 26: Demo!
https://www.youtube.com/watch?v=Cojl8VEp_cA

## Slide 27: Remaining Contemporary Work
Finish web server code that parses web form into C code arguments
Fix communication board
Documentation
Notes: Norm

## Slide 28: Estimated cost
PCBs: ~$220
  Two four-layer boards + shipping: $170
  One two-layer board + shipping: $50
Parts for PCBs: ~$250 (not including ZIF socket)
Numato shield: ~$80 (student discount)
BeagleBone Black: ~$50
Total cost: ~$600 (closer to ~$1000 if we include failed boards and repurchasing parts)
Notes: Norm

## Slide 29: Future Ideas: Web Server
Security measures…
  Anyone can log in as root right now…
User accounts and passwords
Store test vectors in database
Schmooing plots
Notes: Norm

## Slide 30: Future Ideas: Hardware
Bigger, better FPGA
  Xilinx Spartan-6 isn’t designed to handle very high frequencies
  7-series Xilinx FPGAs are designed for fast clock speeds, but are extremely expensive
  Snickerdoodle board: https://www.crowdsupply.com/krtkl/snickerdoodle
    ZYNQ SoC (ARM processor + FPGA)
    Provides 150 I/O pins
    $150
Conditional test vectors may be possible with two FPGAs, or an FPGA with ~300 pins
  Area of further investigation
Notes: Daniel

## Slide 31: Design Challenges: Verilog
Initial Verilog design included R0, R1, DNRZ_L, DNRZ_T force formats (2-bit encoding)
  Design required 108% of available LUTs
Removed DNRZ_T and R1 force formats (1-bit encoding)
  Design required 70% of available LUTs, but 115% of FPGA area
Removed ability to specify two different test cycles, replaced with one test cycle but two leading edges
  Design now utilizes 40% of available LUTs, 70% of available area
Notes: Daniel

## Slide 32: Reflection and Impact
Project pushed us to our limits
  Verilog is over 2000 lines, more than 50 states for “main” FSM
  Hardware boards by far the most complicated we have ever designed
Proud of our design choices
  Things we learned (amongst other things):
    How to utilize hierarchical schematics and autorouting rules in Altium
    The Spartan-6 FPGA is (relatively) limited!
    Glitch-free FPGA outputs can be difficult to achieve
      Double-buffering signals seems to help
    PHP can call C executables!
Notes: Daniel

## Slide 33: Special Thanks
Ken Stevens
  Awesome project advisor, provided useful feedback
Erik Brunvand
  Provided (expensive) ZIF socket, FPGA boards, and useful feedback
Advanced Circuits
  Reasonable prices (and free popcorn with every board purchase)
Notes: Daniel and Norm wrap up talk